**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Ambulatory BP

**Brand** Suntench

**Model** Oscar 2

**Serial/Asset #** 113500

**Department** CPN (HEART CENTER)

**TEST EQUIPMENT**

Biotek Instrument Universal Biometer

**Model:** DPM III

**Serial #:** 113123

# Calibration Verification

**Calibration Date APRIL 18, 2023**

**Due Date APRIL 18, 2023**

**File BTSI-TMCSL-2023-006**

|  |  |  |
| --- | --- | --- |
| **Device Under Test** | **Actual Measured** | **Deviation** |
| 180 mmHg | 182 mmHg | +2 mmHg |
| 160 mmHg | 158 mmHg | -2 mmHg |
| 140 mmHg | 141 mmHg | +1 mmHg |
| 120 mmHg | 119 mmHg | -1 mmHg |

**REMARKS:** Tolerance is ±5 mmhg

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***