**CLIENT INFORMATION**

**Calamba Medical Center**

**Crossing, Calamba City, Laguna**

**EQUIPMENT** Analytical Balance

**Brand** Genesis

**Model** CB 220

**Serial/Asset #** 3243

**Department** Laboratory

**TEST EQUIPMENT**

Metler Toledo,GmbH

Weights: 10g,20g,30g,40g,50g

# Calibration Verification

**Calibration Date** May 2, 2022

**Due Date** May 2, 2023

**File** BTSI-CMC-2022-2896

|  |  |  |
| --- | --- | --- |
| **Weights Value** | **As Found** | **Deviation** |
| 10g | 9.99g | 0.01g |
| 20g | 20.01g | 0.01g |
| 30g | 29.98g | 0.02g |
| 40g | 39.99g | 0.01g |
| 50g | 49.98g | 0.02g |

**REMARKS:** Tolerance is ±1%

**OVERALL RESULT: PASS**

Biomedical Technician: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***