**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Anesthesia Machine

**Brand** GE Datex Ohmeda

**Model** M-904E

**Serial/Asset #** MA12120619

**Department** Operating Room

**TEST EQUIPMENT/S**

|  |  |  |  |
| --- | --- | --- | --- |
| **BRAND:** | Datrend | **BRAND:** | Datrend |
| **MODEL:** | VenTest | **MODEL:** | Flow Sensor |
| **SERIAL #:** | DVT1625128 | **SERIAL #:** | FSH1615119 |

# Calibration Verification

|  |  |
| --- | --- |
| **Calibration Date:** | Mar. 15, 2022 |
| **Due Date:** | Mar. 15, 2023 |
| **Control #:** | BTSI-TMCSL-2022-1712 |

**Detailed Test Data**

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| |  |  | | --- | --- | | PAW | | | DUT | 20 cmH2O | | MV | 20 cmH2O | | |  |  | | --- | --- | | Breath Rate I:E | | | DUT | 1:2 | | MV | 1:2 | |
|  |  |
| |  |  | | --- | --- | | Flow Rate | | | DUT | 5 psi | | MV | 4.9 psi | | |  |  | | --- | --- | | Tidal Volume | | | DUT | 500 mL | | MV | 498 mL | |

**REMARKS: The unit is in good condition and acceptable to use.**

**OVERALL RESULT: PASS**

|  |  |
| --- | --- |
|  |  |
| **Biomed Technician**: Cristopher M. Nipal | **Verified By:** Engr. Oliver E. Sta. Ana |

**\*\*end of record\*\***