**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Centrifuge

**Brand** Hettich

**Model** Haematokrit 210

**Serial/Asset #** D78532

**Department** Laboratory

**TEST EQUIPMENT**

**Digital Tachometer**

**Model:** DT-2236C

**Serial #:** S338496

# Calibration Verification

**Calibration Date** Mar. 17, 2022

**Due Date** Mar. 17, 2023

**File** BTSI-TMCSL-2022-2400

|  |  |  |
| --- | --- | --- |
| **Device Under Test** | **Actual Measured** | **Deviation** |
| 3000 | 3018 RPM | +18 |
|  |  |  |
|  |  |  |
|  |  |  |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***