**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Baby Scale

**Brand** EBSA

**Model** Ebsa-20

**Serial/Asset #**

**Department** NICU

**TEST EQUIPMENT**

Metler Toledo,GmbH

Weights: 1kg, 2kg, 5kg,10kg

# Calibration Verification

**Calibration Date** June 21, 2023

**Due Date** June 21, 2024

**File** BTSI-TDMH-2023-041

|  |  |  |
| --- | --- | --- |
| **Weights Value** | **As Found** | **Deviation** |
| 1kg | 1kg | 0kg |
| 2kg | 2kg | 0kg |
| 5kg | 5kg | 0kg |
| 10kg | 10kg | 0kg |
| 15kg | 15kg | 0kg |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***