**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** ECG Machine

**Brand** GE

**Model** MAC i

**Serial/Asset #** SP01112077IPA

**Department** Dialysis

**TEST EQUIPMENT**

**Datrend Electrical Safety Analyzer**

**Model:** vPad Rugged 2

**Serial #:** VPR17030267

# Calibration Verification

**Calibration Date** Mar. 16, 2022

**Due Date** Mar. 16, 2023

**File** BTSI-TMCSL-2022-1732

**Detailed Test Data**

|  |  |
| --- | --- |
| ECG Test | |
| **Target Value** | **Measured Value** |
| 30BPM | 30BPM |
| 60BPM | 60BPM |
| 120BPM | 120BPM |
| 180BPM | 180BPM |
| 240BPM | 240BPM |
| 300BPM | 300BPM |
| Performance Test | |
| **Test Type** | **Result** |
| Sinus Rhythm | Good |
| Square Wave | Good |
| Triangle Wave | Good |
| Sine Wave | Good |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***