**CLIENT INFORMATION**

**Taytay Doctors Multispeciality Hospital, Inc.**

**Rizal Avenue Taytay, Rizal**

**EQUIPMENT:** Fetal Monitor

**Brand: China**

**Model: N/A**

**Serial/Asset #:** N/A

**Department: Operating Room**

**TEST EQUIPMENT**

**Datrend Vital Sign Simulator Datrend Electrical Safety Analyzer**

**Model:** vPad-PS **Model:** vPad Rugged 2

**Serial #:** VSM17020082 **Serial #:** VPR17030267

# Calibration Verification

**Calibration Date: February 23, 2023**

**Due Date: February 23, 2024**

**File**

**Detailed Test Data**

|  |  |
| --- | --- |
| Heart Rate Test | |
| Target Value | Measured Value |
| 30 BPM | 33 BPM |
| 60 BPM | 61 BPM |
| 120 BPM | 119 BPM |
| 180 BPM | 176 BPM |
| 240 BPM | 244 BPM |
| 300 BPM | 299 BPM |

**REMARKS:**

**OVERALL RESULT: Pass**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***