**CLIENT INFORMATION**

**PLTC Inc. LUIS A. Tiam Medical Center**

**National Highway, Brgy. Bascaran, Solano, Nueva Vizcaya**

**EQUIPMENT:** Fetal Doppler

**Brand:** Hello Baby

**Model:** SC-307

**Serial/Asset #:** 7034E100198010100360

**Department: Delivery Room**

**TEST EQUIPMENT**

**Datrend Vital Sign Simulator Datrend Electrical Safety Analyzer**

**Model:** vPad-PS **Model:** vPad Rugged 2

**Serial #:** VSM17020082 **Serial #:** VPR17030267

# Calibration Verification

**Calibration Date:**

**Due Date:**

**File**

**Detailed Test Data**

|  |  |
| --- | --- |
| Heart Rate Test | |
| Target Value | Measured Value |
| 30BPM | 30BPM |
| 60BPM | 60BPM |
| 120BPM | 120BPM |
| 180BPM | 180BPM |
| 240BPM | 240BPM |
| 300BPM | 300BPM |

**REMARKS:**

**OVERALL RESULT:**

Biomedical Technician: Engr. Oliver E. Sta. Ana Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***