**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Infant Incubator

**Brand** Isolette

**Model** Airshields C100

**Serial/Asset #** HK10610

**Department** NICU

**TEST EQUIPMENT**

Biotek Instrument Universal Biometer

**Model:** DPM III

**Serial #:** 113123

# Calibration Verification

**Calibration Date** Mar. 18, 2022

**Due Date** Mar. 18, 2023

**File** BTSI-TMCSL-2022-1702

|  |  |  |
| --- | --- | --- |
| **Device Under Test** | **Actual Measured** | **Deviation** |
| 35.8°C | 36.2°C | 0.4°C |
|  |  |  |
|  |  |  |
|  |  |  |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***