**CLIENT INFORMATION**

**St. John the Baptist Medical Center INC.**

**Brgy. Parian, Calamba, Laguna**

**EQUIPMENT:** Infusion Pump

**Brand: Szosen**

**Model:** OIP-900

**Serial/Asset #: 002109711**

**Department: NICU**

**TEST EQUIPMENT**

**Datrend Infusion Device Analyzer**

**Model:** Infutest 2000

**Serial #:** VPR17030267/EDC17031522

# Calibration Verification

**Calibration Date: February 22, 2023**

**Due Date: February 22, 2024**

**File**

**Detailed Test Data**

|  |  |  |  |
| --- | --- | --- | --- |
| **Target Avrg.(mL/h)** | **Actual Measured(mL/h)** | **Time(min)** | **Volume Infused(mL)** |
| 100 | 100.6 | 3:12 | 5.0mL(5.2mL) |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |

**REMARKS:**

**OVERALL RESULT: Pass**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***