**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Patient Monitor

**Brand** GE

**Model** Dash 4000

**Serial/Asset #** SD010397953GA

**Department** OR

**TEST EQUIPMENT**

**Datrend Vital Sign Simulator Datrend Electrical Safety Analyzer**

**Model:** vPad-PS **Model:** vPad Rugged 2

**Serial #:** VSM17020082 **Serial #:** VPR17030267

# Calibration Verification

**Calibration Date** Mar. 15, 2022

**Due Date** Mar. 15, 2023

**File** BTSI-TMCSL-2022-1714

**Detailed Test Data**

|  |  |  |  |
| --- | --- | --- | --- |
| ECG Test | | NIBP Test | |
| Target Value | Measured Value | Target Value | Measured Value |
| 30BPM | 30BPM | 120-80-93 | 123-78-93 |
| 60BPM | 60BPM |  | |
| 120BPM | 120BPM | SPo2 Test | |
| 180BPM | 180BPM | Target Value | Measured Value |
| 240BPM | 240BPM | 99% | 98% |
| 300BPM | 300BPM |  | |
| Performance Test | | Respiration Rate Test | |
| Sinus Rhythm | Good | Target Value | Measured Value |
| Square Wave | Good | 20 | 20 |
| Triangle Wave | Good |  |  |
| Sine Wave | Good |  |  |
|  |  |  |  |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***