**CLIENT INFORMATION**

**Calamba Medical Center**

**CMC Avenue, Brgy. Real Calamba, Laguna**

**EQUIPMENT** Pipette

**Brand DRAGON LAB**

**Model** 10 - 100 ul

**Serial/Asset #** YL20AAK0011844

**Department** MOLECULAR

**TEST EQUIPMENT**

Electronic Balance

500g/0.001g

# Calibration Verification

**Calibration Date**: OCTOBER 14, 2022

**Due Date:** OCTOBER 14, 2023

**File:** 3793

|  |  |  |
| --- | --- | --- |
| **Set Value** | **As Found ( 1mL/g )** | **Deviation %** |
| 10 ul | 11 g | -1% |
| 40 ul | 39 g | +1% |
| 60 ul | 60 g | 0% |
| 100 ul | 100 g | 0% |
|  |  |  |

**REMARKS:**

**OVERALL RESULT: PASSED**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***