**CLIENT INFORMATION**

**Taytay Doctors Multispeciality Hospital, Inc.**

**Rizal Avenue Taytay, Rizal**

**EQUIPMENT** Pulse Oximeter

**Brand** Nellcor

**Model** Covidien

**Serial/Asset #** MBB1508087

**Department** NICU

**TEST EQUIPMENT**

**Datrend Vital Sign Simulator**

**Model:** vPad-PS

**Serial #:** VSM17020082

# Calibration Verification

**Calibration Date** Mar. 18, 2022

**Due Date** Mar. 18, 2023

**File** BTSI-TMCSL-2022-1757

**Detailed Test Data**

|  |  |
| --- | --- |
| **Target(O² Sat. & BPM)** | **Actual Measured** |
| 100% | 99% |
| 60 | 60 |
| 70 | 70 |
| 80 | 80 |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***