**CLIENT INFORMATION**

**Taytay Doctors Multispecialty Hospital, Inc.**

**Rizal Avenue Taytay, Rizal**

**EQUIPMENT** Sphygmomanometer

**Brand** W.A. Baum

**Model** Floorlamp

**Serial/Asset #** 06083036

**Department** OR

**TEST EQUIPMENT**

Biotek Instrument Universal Biometer

**Model:** DPM III

**Serial #:** 113123

# Calibration Verification

**Calibration Date June 22, 2023**

**Due Date June 22, 2024**

**File BTSI-TDMH-2023-008**

|  |  |  |
| --- | --- | --- |
| **Device Under Test** | **Actual Measured** | **Deviation** |
| 180 mmhg | 178 mmhg | -2 mmhg |
| 160 mmhg | 158 mmhg | -2 mmhg |
| 140 mmhg | 138 mmhg | -2 mmhg |
| 120 mmhg | 118 mmhg | -2 mmhg |

**REMARKS:** Tolerance is ±5 mmhg

**OVERALL RESULT: PASS**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***