**CLIENT INFORMATION**

**The Medical City South Luzon**

**Greenfield City, Sta. Rosa City, Laguna**

**EQUIPMENT** Syringe Pump

**Brand** B. Braun

**Model** 8713030

**Serial/Asset #** 316732

**Department** NICU

**TEST EQUIPMENT**

**Datrend Infusion Device Analyzer**

**Model:** Infutest 2000

**Serial #:** VPR17030267/EDC17031522

# Calibration Verification

**Calibration Date** Mar. 18, 2022

**Due Date** Mar. 18, 2023

**File** BTSI-TMCSL-2022-1751

**Detailed Test Data**

|  |  |  |  |
| --- | --- | --- | --- |
| **Target Avrg.(mL/h)** | **Actual Measured(mL/h)** | **Time(min)** | **Volume Infused(mL)** |
| 900 | 840.6 | 3:08 | 43.9 |
| 300 | 287.0 | 3:06 | 14.83 |
| 30 | 28.26 | 3:18 | 1.55 |
|  |  |  |  |

**REMARKS:**

**OVERALL RESULT: PASS**

Biomedical Technician: Cristopher M. Nipal Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***