**CLIENT INFORMATION**

**EQUIPMENT**

**Brand** Safeplus

**Model**

**Serial/Asset #** H6E19M003730

**Department** Nurse Station 4B

**TEST EQUIPMENT**

Biotek Instrument Universal Biometer

**Model:** DPM III

**Serial #:** 113123

# Calibration Verification

**Calibration Date June 22, 2023**

**Due Date June 22, 2024**

**File BTSI-TDMH-2023-017**

|  |  |  |
| --- | --- | --- |
| **Device Under Test** | **Actual Measured** | **Deviation** |
| 37°C | 36.5°C | -0.5°C |
|  |  |  |
|  |  |  |
|  |  |  |

**REMARKS:**

**OVERALL RESULT:**

Biomedical Technician: Christian F. Ynzon Verified By: Engr. Oliver E. Sta. Ana

**\*\*end of record\*\***